## Overview

The Tecan talks to the “Cf.Daq.Server.exe” program using command line arguments sent to the “Cf.Daq.Command.exe” executable:

CfCommand ClearError

These commands return immediately and to see if motion completed, the corresponding status command needs to be sent.  Each command returns an integer representing success/failure or a value indicating the requested state.  Errors require explicit clearing via the ClearError command.

The server can be run in either debug or regular mode depending on whether the Cf.Daq.Controller.dll is present in the server folder.  If it is not there, the Cf.Daq.DummyController.dll mock controller is loaded instead.  The mock controller allows you to run without NI-DAQ being installed and will simulate all operations and their timing.

For testing on the instrument, use the Cf.Daq.ClientGui.exe which will send commands to the server.

## ClearError

Must be used if an error occurs before additional commands can be executed.

Returns 0 on failure, 1 on success.

## ConfigureTecc

Returns  0 on Failure, 1 success.  Configures TECC with current settings in Server UI.

## ConfigureNewTeccFile [File_Name]

Returns  0 on Failure, 1 success.  Creates a new TECC file.  Must be called before a call to TakeTeccSample.  If File_Name is not specified, program will just make a time stamped name.

## GetElectrodeManifoldStatus

Returns  0 on Failure, 1 on Idle, 2 when moving

## GetTeccStatus

Returns  0 on Failure, 1 on Idle, 2 when recording

## SetManifoldPosition Position [Z_Position]

Returns 0 on failure, 1 on success.

Possible Positions:

CellPlate_1

CellPlate_2

CellPlate_3

CellPlate_4

CellPlate_5

CellPlate_6

Wash

Rinse

Offset

You can also call positions by number (0-31).  If Z_Position is not specified, default position is up.

Possible Z_Positions:

Up

Down

## SetServo State

Turns Yamaha servo on/off.  Returns 0 on failure, 1 on success.

Possible States:

On

Off

## SetWash State

Returns 0 on failure, 1 on success.

Possible States:

On

Off

## SetZAxisPosition Position

Returns 0 on failure, 1 on success.

Possible Positions:

Up

Down

## TakeTeccSample Position [Step_Name] [ MeasurementTimeS]

Move Electrode manifold to the specified positions, take a measurement, and wash manifold.  Returns 0 on failure, 1 on success. To see when done, use the GetElectrodeManifoldStatus call.  The Step_Name field is an optional parameter that is automatically put into the output file to identify what is happening where.  If measurement time is specified, then Step_Name is mandatory!

Possible Positions:

CellPlate_1

CellPlate_2

CellPlate_3

CellPlate_4

CellPlate_5

CellPlate_6

Offset

Measurement time in seconds only used for debugging!

## TakeTeccSampleNoMove [ MeasurementTimeS]

Take an in place measurement with no moves. Measurement time in seconds only used for debugging!

## WashAndWait WashCycles

Will move to the wash station and dip the manifold the number of times specified in WashCycles.  Valid values of 0-20.

Returns 0 on failure, 1 on success.