## Proposed Extremely Rough and Unvetted Project Timing

- December 2014

  - Design of deck layout without final manifold adapter and other smaller items

  - Yamaha on order/promise from Japan

- Jan 2014

  - Delivery of TECC-24 and electrode manifold

  - Mechanical design complete and all parts ordered

  - Software testing started with deck mock-up. This is critical because we need to determine whether we can use EVO Ware or need a custom scheduler.

- Feb-Mar 2014

  - Testing/rework with Yamaha

- April 2014

  - Delivery

## System Timing

- 5 minute measurement interval per plate

- Plate Measurement Cycle: (26 s measurement + 8 s move  + 7 s rinse) * 7 = 287 s

  - Treat offset measurement exactly as a normal plate measurement

  - Move and rinse are a guess.  Reducing rinse/move time makes timing less critical.

  - Move time will be position dependent

- If possible, when the number of plates is reduced, the system will decrease the sampling internal to take more data points.

- Full protocol is ~2hr of run time

- Critical that compounds are delivered to plates while data is collected (on separate plates).

## Software Control

- Tecan EvoWare will control scheduling and all operations unless there are issues with scheduling.  In this case, an external program will call routines developed in EVO Ware.

  - Need to make sure Yamaha motions do no conflict with MCA!

  - Fluidic operations need to be very tightly coupled to measurements

- Electrode positioning will be controlled via DIO

- Measurements will be triggered by relay closures and the TECC software configured for each run by the user.

- Goal is to allow user to have enough flexibility to make changes without SPaRCS involvement in long term.

## Temperature Control

- Tecan EVO will be enclosed and heated

- Transwell plates will be heated from below using either Al block with DC blanket heater or with brass block and heated water.

- Electrode Manifold has internal heater  and sensor (controller was ordered from Willy)

- Where do temp sensors go for internal air and Transwell heater?

## Electrode Manifold Movement

- Yamaha X-Axis with 1250 mm stroke will move between the 6 Transwell locations, rinse plate, offset plate.  Only ~1000 mm of stroke needed with close plate spacing.

- Air cylinder (Z-Axis) will provide 1” of stroke

- Z-Axis sensors for up/down position with the X-Axis motion being locked (with manual switch override) automatically when electrodes are in down position.

- Cable track setup?

- IDC ribbon connectors coming out of manifold.  Willy seems unwilling to send us a solid model.

## Deck Components 24-Well System

- 6 Costar 24-well Transwell Plates in a line accessible with MCA + 2 plates positions for wash and offset measurement that need to be accessible to Yamaha.

  - Cleaner to put in back of instrument with Yamaha on deck, but restricts access to plate hotels

  - Put in front with Yamaha off deck increases usable space on deck but makes instrument deeper.

  - Transwell plates do not need covers.

- Enough 96 position tip boxes for 8 distinct sets of dispense/mix cycles: 8X24/96 = 2

  - These are custom 96 position tip boxes with a English offset to match MCA head and Costar format

  - Automated tip racker?

  - Tip box design?

- 8 24-well plates for reagents

  - Stored in plate “hotels” accessed by Tecan gripper (TeMo)

    - How many “active” reagents plate sites can be fit on deck?

    - How long does it take to replace one?

    - Time to move plate from hotel to deck?

- Active wash reservoir for rapid clean of electrode tips.  Ideally just a few seconds of dip time.

  - Streaming?  Stirred? Ultrasonic? Air/Buffer combo?

  - Needs to be heated

  - Does not need to be exchanged during the run

- Offset Measurement Plate

  - Just a normal plate position

  - Middle or at the ends of the X-Axis.

## Deck Components 96-Well System Upgrade mid/late-2014

- Upgrade to 6 Costar 96-well Transwell Plates

- Enough 96 tip boxes for 8 distinct sets of dispense/mix cycles: 8

  - Hoteling or off-deck stacker?

- 8 96-well plates for reagents

# Critical Questions

- Position of Yamaha (X-Axis).  Front or back?  On/Off deck?

- Will alignment of plates to the X/Z stages be good enough across 1 m?  We have 0.037” slop in X and 0.060” slop in Y.

- Is Z height easy to adjust with new plate lots?

- Elevator size in Bob’s building

- Do we have enough deck space for 96 well plate upgrade? If not, can we locate resources off deck?